## Procurement Plan
Annex A - Specifications List for Medical Equipment
ITB-HCR-ROK-2022-004
Supply and Delivery of Medical Equipment for UNHCR Sudan Operation
# | Name of the items | Specification | Quantity
1 | Glucomachine Kit - The machine and strips should be compatible | The Glucometer blood glucose meter is equipped with many powerful features into a smaller size. No coding reduces handling steps, see test result averages up to 90 days, get alerts if your test strips are expired or if you've underdosed, re-dose within 10 seconds, set pre-and post-meal reminders to test, and more. he device allows for easy and accurate measurement. The kit contains: 1 Glucometer, 1 Softclix Lancing Device, 10 Softclix Lancets, 50 Active Test Strips, User Manual and Carrying Case. Technical Specifications: Measuring interval: 10-600 mg/dL, Blood volume: 1 - 2 microlitre, Measuring time: Test strip in the meter: approx. 5 seconds. Test strip outside the meter: approx. 08 seconds, Power supply: 1 battery (type CR2032), Battery life: approx. 1000 tests or approx. 1 year, Automatic power-off: after 30 or 90 seconds depending on operating status, Memory: 500 test results with time and date, as well as 7, 14, 30 and 90 days averages | 20 Kit
2 | Glucostrip - The strips and the machine should be compatible | Strips for blood Glucometer (ex: Accu-Chek® ), Model - Strips Active Test Strips - 50, Minimum Blood Sample Needed -1.2 microlitre, Measuring Range- 10 mg/dL - 600 mg/dL, Number of Strips- 50 Strips, Number of Lancets- 50 Lancets, Sales Package- 1 Vial of 50 Test Strips, 1 Code Chip, Other Features: Re-dosing Capability, Visual Color Check : Plausibility Check Option with 5 Color Block on the Test Strip Vial | 20,000 Strips
3 | Lancets | Lancet,blood,ster,disp/PAC-200. Lancet, blood, sterile, disposable, pack of 200.Typical incision depth, approximately 2mm Single use Sterile individually packed Material: Chrome-polished stainless steel Dimensions: Length, approximately 40mm Width, approximately 6mm Thickness, approximately 0.05mm | 100 Pcs
4 | ECG machine with Accessassories - please ensure ECG and rollls are compatible | General description: Electrocardiograph (ECG) digital monitor and recorder, 12-leads detection, multi-channel recording, portable, AC and battery powered, with printer and accessories. Intended Use: Electrocardiographs detect the electrical signals associated with cardiac activity and produce an ECG, a graphic record of the voltage versus time. They are used to diagnose and assist in treating some types of heart disease and arrhythmias. Electrocardiographs differ from ECG monitors, which provide a continuous real-time display of the ECG waveform and other physiologic measurements on a screen and warn of potentially life-threatening or precursor rhythms. Electrocardiographs are used to diagnose cardiac abnormalities, determine a patient's response to drug therapy, and reveal trends or changes in heart function. Multichannel electrocardiographs record and display signals from two or more leads simultaneously. TECHNICAL SPECIFICATIONS: ECG analysis and full interpretation (rhythm and events), real-time and manual interpretation. Three (3) or more recording channels. SIGNAL ACQUISITION: Simultaneous 12 lead acquisition ( aVR, aVL, aVF, I, II, III and V1 - V6) derived from 10 electrodes (RA, LA, RL, LL, V1 - V6). Arrythmia and ST elevation detection. Minimum gain/sensitivity settings include 2.5, 5, 10 and 20 mm/mV. Lead switching, either manual, automatic or both (manual and auto options). Adjustable trace speeds include 5, 10 (and/or 12.5), 25 and 50 mm/s. Minimum HR range 30 – 300 bpm with rhythm analysis. Minimum guaranteed diagnostic frequency response of 0.05 – 150 Hz. Common Mode Rejection Ratio (CMRR) at 60 Hz > 105 dB or better. Calibration signal of 1 mV, manual and/or automatic. Selectable/adjustable filters for baseline drift, muscle artefacts, mains power. Pacemaker detection. Accuracy of input signal reproduction ± 5 % or ± 40 μV, whichever is greater. Input impedance > 50 MΩ. Internal noise level < 12.5 µV peak-to-peak. Automatic baseline centring. Defibrillation fluctuation/overload protection. Baseline recovery < 5 s after defibrillation. AC fluctuation protection. RECORDER AND PRINTER: Minimum of 3 recording channels. Recorder display includes date/time, patient data and heart rate and basic settings. Capable of displaying one group of at least three channels simultaneously. Recorder waveform display includes lead marker and timing marker. Integrated/built-in printer. Capable of printing user selected number of channels. Capable of printing one group of at least three channels simultaneously. Paper speeds include 5, 10 (or 12.5), 25 and 50 mm/s. Compatible with Z-fold paper and optionally with roll paper also (indicate compatibility). DATA INPUT/OUTPUT, STORAGE AND ALARMS : Integrated alpha-numeric keyboard. Patient data input fields include name, age, height and weight, gender. Backlit Liquid Crystal Display (LCD) display screen, minimum of 12.7 cm (5 inches) diagonally. Capable of internally storing a minimum of 10 waveforms for later retrieval, printing and/or transmission. Expandable storage for additional waveforms if required, via USB. Capable of exporting waveform data and reports, via USB or LAN. Visual alarms for patient connection (lead faults), heart rate, printer and paper errors, and system errors. Automatic self-test at start up. POWER PROVISIONS : Built-in rechargeable lithium-ion battery. Minimum battery operating time is 100 ECG exams or 4 hours of continuous recording. Automatic switch to battery in case of power failure, automatic recharge on connection to mains. Maximum battery charging time to full charge is 8 hours. Power requirements: 100 - 240 Volts - 50/60 Hz. Designed for frequent and easy dismount and disinfection with hospital-grade products. SUPPLIED WITH: Instructions for assembly, use and maintenance in English, French and Spanish. 1 x Plastic protective dustcover. 1 x Patient cable. 2 x full sets of chest electrodes. 2 x full sets of limb electrodes. 1 x supply of 960 Z-folded sheets suitable for the unit. (if accepts Z-folded sheets), or 5 x supply of rolls suitable for the unit (if the unit accepts rolls). 3 x bottle of 100 ml of ECG conductive gel. 1 x spare rechargeable battery pack. ESTIMATED LIFE SPAN: Five years. WARRANTY: Two years from shipping date. ENVIRONMENTAL CONDITIONS : Operating conditions: 5°C – 40°C / 25% - 85% RH. Storage conditions: -20°C - 40°C / 25% - 95% RH. Atmospheric pressure: 860 ~ 1060 hPa. WEIGHT AND VOLUME: Weight: 10.00 kg. Volume: 80.00 dm³. MAINTENANCE/USER REQUIREMENTS: As per user and service manuals. | 8 Machine
5 | ECG Paper - please ensure ECG and rollls are compatible | Material: Thermal, Number of Sheets Per Pack: 10, GSM: 80 -120, Application: ECG Machine, Packing type: roll, Features : Smooth texture, Tear resistance, Waterproof, Uses: Hospitals, Clinics, Diagnostic centers | 768 Rolls
6 | CBC machine | o Throughput: 80 tests/hour o Parameters: 22 hematology parameters, including 3-part differential: WBC, LYM, MID, GRA, LYM%, MID%, GRA%, RBC, HCT, MCV, RDW, HGB, MCH, MCHC, PLT, PCT, MPV, P-LCC, P-LCR, RDW-CV, RDW-SD, PDW-CV, PDW-SD o Volumetric impedance method o Sample volume: 25 μl of whole blood in normal 3-part mode, 50 μl of pre-diluted blood o Sampling method: Open tube system with automatic sample rotor o 3 cyanide-free reagents: Diluent, Lyse, Cleaner o Diatro•ECO Reagent Pack: Up to 145 measurements o Aperture diameter: 70 μ (RBC/PLT); 100 μ (WBC) o Clog prevention: High-voltage burst on aperture in each analysis cycle and high pressure back-flush o Data capacity: 10 000 results, including RBC, PLT, and WBC 3-part histogram o Operating temperature: 15—30 °C (59—86 °F). Optimal temperature is 25 °C (77 °F) o User interface: Easy-to-use, menu driven with touchscreen, separate START button and status LED o User languages: English, Indonesian, German, Italian, French, Czech Portuguese, Polish, Spanish, Vietnamese, Russian, Greek, Romanian, Hungarian o Size (Width x Depth x Height): 33 cm x 31 cm x 38 cm, 13” x 10” x 15” o Weight: 12 kg, 26 lbs | 10 Machine
7 | Echocardiogram | 1. Should be a stand alone system integrated on a light weight mobile cart. 2. The system should be a color Doppler Echocardiography all digital beam former system to study the anatomical abnormalities and blood flow in the heart and associated vessels. Should be a stand alone system integrated on a light weight mobile cart. 3. Should be a latest generation Electronic Phased array Color Doppler system with minimum 512 Electronic independent channels. 4. Should have 256 gray shades for sharp contrast resolutions. 5. Should be supplied with adult and pediatric cardiac and vascular probes of wide band transducers without frequency selection for higher sensitivity of response over a broad frequency range of operation. 6. Should have 2D, M-mode, Anatomical M-mode, Color M-mode, PW and CW doppler, Steerable CW doppler. 7. The system should have a very high dynamic range of at least 200dB to pick up subtle echoes. 8. Should have three active ports. 9. Should have 2-4 Mhz broadband phased array sector probe for adult cardiac imaging. 10. Should have 3-8 Mhz broadband phased array sector probe for paediatric and neonatal cardiac imaging. 11. Should have 3-12 Mhz broadband Linear Array probe for vascular imaging. 12. Should have multi frequency convex array probe 3-10 MHz for paediatric imaging 13. Pencil probe (optional) 14. Should have advanced tissue Harmonic Imaging. 15. Should have color flow imaging. 16. Should have color Tissue Doppler Imaging. 17. Should have gain control in Axial Plane. 18. Should have triple imaging possibility on the system. 19. Should have PW/CW Doppler facility in all imaging phased array sector probes. 20. Should have 15” or more high resolution TFT monitor with tilt and swivel facility and should be able to view in all angles and all light conditions. 21. Should have greater than 5000 images in the system hard disk drive 22. Should have in built CD/DVD writer. 23. Should have patient reporting page with embedded images. 24. Should have full functional measurement facility and calculation should be possible. 25. Should be supplied with thermal printer and 6 packs of thermal paper and the unit should have option to connect external printer. 26. Unit should function with 200-240Vac, 50/60 Hz input power supply. 27. Unit should be supplied with suitable UPS with a minimum 30 minutes back-up time. | 8 Machine
8 | Adult Stethoscope | Adult diaphragm Ø: 45,5mm; . Tube made of PVC and is crack resistant. Tube impervious to outside noises, guaranteeing full transmission of sound, good auditive quality. Tube diameter: outer diameter 10mm, inner diameter 4.8mm. | 100 Pcs
9 | Pediatric Stethoscope | paediatric diaphragm Ø: 31.5mm, Tube made of PVC and is crack resistant. Tube impervious to outside noises, guaranteeing full transmission of sound, good auditive quality. Tube diameter: outer diameter 10mm, inner diameter 4.8mm. | 100 Pcs
10 | Sphygmomanometer | Sphygmomanometer, anaeroid, adult | 100 Pcs
11 | Sphygmomanometer | Sphygmomanometer, anaeroid, paediatric | 100 Pcs
12 | Thermometer | Thermometer, clinical, digital, 32-43 Celsius, C/F | 100 Pcs
13 | Torch | Lighting Color, Yellow ; Body Material, Aluminum | 150 Pcs
14 | Weighing Scale | Mechanical Adult Scale, beam type, for use up to 180kg with 100g precision | 40 Pcs
15 | WeighingScale | Scale, electronic, mother-and-child, 150 kg x 100 g , UNISCALE, Type Seca. | 40 Pcs